INTRA-X Mk IV

Automatic 
X-ray Film 
Processor


Installation, operation and maintenance manual

WARNING:

Do not plug in to electrical mains power before reading this manual -or before filling the machine with liquids

Machine serial number to be quoted on all correspondence

MEDIVANCE INSTRUMENTS LIMITED, BARRETTS GREEN ROAD. HARLESDEN, LONDON NW10 7AP TELEPHONE: +44 (0)20 8965 2913

Table of Symbols

| Symbol | Description |
| --- | --- |
|  | Alternating current |
|  | Off (Power: disconnection from the mains) |
|  | On (Power: connection to the mains) |
|  | Type B equipment |
|  | Film transport switch/Process 'RUN' Button |
|  | Attention, consult accompanying documents |

Some machines are fitted with an ON/OFF switch illuminated RED. When lit this signals that the machine is switched ON. In darkroom conditions, red Illumination is used to prevent film fog, and on this machine it is not indicating a dangerous condition.

MAINS POWER SUPPLY:

230 Volts, 50 Hz, 13 Amps, 3-wlre earthed socket. Use only with the connection lead supplied, fitted with a 10 Amp fuse.

Contents

INTRODUCTION TO THE PROCESSOR

The diagrams at the front of this manual are designed to introduce you visually to the layout of the Velopex.

Page

EXTERNAL FEATURES OF THE MACHINE                      						3 
INTERNAL FEATURES OF THE MACHINE                      						4

SITING OF THE VELOPEX

Pre-lnstallation Instructions                                                     						5 
Utilities                                                                          							5

PREPARING TO USE THE MACHINE

Unpacking                                                        								6 
Filling the Machine with Chemicals                                  						6 
Daylight Loader                                                    							7

OPERATING THE VELOPEX

To Start                                                             								8 
Processing Films                                                      							8 
Endo Interrupt                                                        							8 
Information Panel                                                     							9 
Film Collector                                                         							9 
Each Day                                                            								9

ROUTINE MAINTENANCE

Cleaning the Processor                                               						           10 
Caution                                                             							           11

TROUBLESHOOTING

Symptoms and Recommended Action                             						           12 
Gear Replacement                                            							           13

FILM QUALITY PROBLEMS

Films too dark                                                        						           14 
Films too light                                                        						           14 
Films dirty or marked                                                  						           15

CONTAMINATION                                           							           16

WARNING: STUDY THIS MANUAL THOROUGHLY BEFORE ATTEMPTING TO INSTALL OR OPERATE PROCESSOR. FAILURE TO DO SO COULD RESULT IN IMPROPER OPERATION WHICH MAY LEAD TO INJURY OR MACHINE DAMAGE.

External Features of the Machine

Optional rigid daylight loader

1 Control Panel

2 On/Off Power Switch

3 Ready Light

4 Fuse Holder

5 Endo Interrupt Facility

6 Air Vents

7 Daylight Loader

8 'RUN' Button

9 Machine Lid

10 Protruding Grip

11 Film Collector

12 Double-fused Chassis Plug

13 Film Entry Guide

14 Hand Entry Ports

15 Viewing Window

Internal Features of the Machine

1	Module Turning Tool 
2	Transport Module Gears 
3	Synthetic Film Transport Belt 
4	Endo Interrupt Facility 
5	Drive Shaft 
6	Drive Dog

7	Developer Tank and Module 
8	Fix Tank

9	Water Tank 
10	Dryer Module 
11	Solution Heaters 
12	Paddle

13	Paddle Drive 
14	Dryer Vent

15	Temperature Sensor 
16	Level Sight Glass

Note

1 1/2 quarts (1400ml) of chemical without module in place will bring tank to correct liquid level. Placing the module in the tank displaces the liquid bringing it up to working level (to middle point of sight glass). This level should be maintained throughout the working life of the chemical.

Siting of the Velopex

PRE-INSTALLATION INSTRUCTIONS

Although the Velopex is designed for use in daylight, care should be taken to avoid sources of intense light.

For example do not site directly under a window, fluorescent tube or floodlight. A well ventilated position is mandatory.

The Intra-X weighs 31 Ib (141<g) when filled; a stable and level counter top is required, sufficient to support weights up to at least 100 Ib (50kg). It is important that when the machine is filled with solution the stand does not rock or shake, as this could cause spillage. NEVER move the machine with chemicals in the tanks.

Avoid siting the machine above or near other electrical, mechanical or sensitive apparatus - e.g. autoclave, compressor, evacuation equipment.

When the machine is to be sited in a darkroom, it is essential to ensure that the darkroom is well ventilated and light-tight, and that there is sufficient space for easy chemical changes and for the machine to be serviced safely. When choosing a site for your Velopex it may be convenient to site the machine near a sink for ease of changing chemicals.

The machine must never be sited below surfaces that could restrict air circulation around it (e.g. below or inside a cupboard). In choosing the site, it is important to allow for the presence of liquids: surfaces susceptible to water (or chemical) damage should therefore be avoided. Avoid carpeted areas.

The ambient temperature must be maintained below 80 degrees (F), 27 degrees (C).

UTILITIES - Electrical

The main power supply must be within 3 feet of the machine and easily accessible for maintenance routines. Electrical lines must be "unitary", i.e. without extensions to the length.

The machine will have been supplied with a mains power cord with moulded plug.

See inside front cover.

WARNING: X-ray radiation can be harmful to patient, technician and dentist. Inadequate lead shielding of the darkroom or film storage area will also cause fogging from exposure of films to stray x-ray radiation. Consult your local codes. Health Department or Dental Equipment Dealer for proper construction of darkroom or placement of film processing equipment in the vicinity of x-ray radiation sources.

Preparing to Use the Machine

UNPACKING

NOTE: For unpacking and lifting the machine into position it is important to have assistance.

The machine comes in a single carton containing: Machine in "darkroom" or "daylight" configuration. Operator's Manual, Film Catcher, Electrical Cord, Transport Module Turning Tool, Cleaning Brush, Chemical Change Chart.

1. Familiarise yourself with the layout of the machine by referring to illustrations at the front of this manual. It is useful to refer to these illustrations as you progress through the manual.

2. Lift the machine from the carton and position on counter top. Remove outer and inner packaging, including carry-straps. The transport modules are protected by internal packing pieces: these must be discarded.

FILLING THE MACHINE WITH CHEMICALS

1. It is important to run the machine with clean water only in the three

liquid-containing tanks with transport modules in position - for a complete running cycle, before starting to fill with chemical. This will clear any dust or debris that may have appeared in the machine during shipment and handling.

WARNING: Ensure that the machine is disconnected from mains power supply whilst filling with chemicals and water. NEVER turn on the machine with the tanks empty.

Do not use chemistry or film designed for manual processing.

NOTE: optimum processor performance will be achieved by using ONLY Velopex Fixer and Developer. The machine is built, however, to work satisfactorily using most high quality brands of chemical.

Always use chemical designed for automatic processing ONLY

Chemicals MUST be from universally accepted manufacturers and be clear, contaminant-free and within expiry date.

2. Check that the film interrupt slide on the lid is in the 'up' position. Open the lid gently by means of the protruding grip, and bring it to the upright position where it will rest. The lid may be completely removed if this is more convenient.

3. Open the machine lid and remove all packing materials.

4. To gain access to the tanks remove the side panel by sliding it upwards in its runners.

5. Remove the dryer module first, situated at the back of the machine, followed by the water tank.

Preparing to use the Machine

6. Remove fixer tank next: NOTE: THIS MUST BE WITHDRAWN BY

SLIDING SIDEWAYS UNTIL COMPLETELY CLEAR OF ITS IMMERSION HEATER MOUNTED ON THE MACHINE INTERNAL WALL. (See diagram, page 4, item 11.)

(Developer tank is removed in the same way.)

7. The transport modules should now be lifted from their respective tanks.

8. Fill each tank with the appropriate chemical, pouring in \Vi quarts (1400ml) of fluid.

NOTE: The Developer and Fixer Tanks have a Level Sight Glass fitted. The correct liquid level will show half way up the sight glass with module in position, and should be maintained at this level at all times.

CONSULT THE CHEMICAL
MANUFACTURER'S SAFETY SHEETS
BEFORE HANDLING THE CHEMICALS.

9. Grease immersion heaters for developer and fixer tanks with the lubricant provided.

10. Carefully lower each transport module into its respective tank - as rapid

immersion will cause excessive displacement of chemical leading to spillage. Should any spillage occur it should be cleaned up immediately.

11. Return tanks to the machine.

12. Re-check the level of the liquids, and top up to centre of sight glass (see diagram, page 4).

13. Engage the module drive pins into the drive dogs, working from right to left (see diagram, page 4).

14. Replace side panel and lid.

15. The machine may now be switched on.

IT WILL NOT OPERATE WITH THE LID OPEN

16. Check that all modules are engaged and running smoothly by passing a "clean-up film" through the machine.

17. If "clean-up film" fails to exit, check that the transport modules are fully engaged in their drive mechanism.

DAYLIGHT LOADER

See diagram, page 3. When the machine is fitted with a daylight loader familiarise yourself with the location of the film entry slot and the process 'RUN' button.

When processing film do not remove hands from the loading section until the film has fully entered the machine. Keep the daylight loader clear of old film wrappers to avoid these entering the machine and to maximise the space available.

Operating the Velopex

TO START

Turn the Main Power Switch ON.

The machine will now run for approximately 8 minutes and begin to bring the chemicals to the correct operating temperature, after which the transport mechanism will stop, placing the machine in 'standby' mode.

The time taken to achieve the correct temperature for processing depends on the room temperature. After reaching the required operating temperature the machine is ready to accept films (average warm-up time 10-15 minutes).

The automatic temperature control will maintain the correct operating temperature for processing: correct temperature is indicated when the temperature indicator light is extinguished. The machine remains in 'standby' mode throughout the day, ready for immediate use.

PROCESSING FILMS

NOTE: The machine should be used only with clean, dust-free films from world-wide recognised manufacturers, and be within expiry date.

In daylight use, put hands through the entry ports in the daylight loader, press the process 'RUN' button on the right hand side of the film entry guide, strip the wrapping from the film with the aid of the viewing window and post the film centrally into film entry slot. Always remove film wrappers from daylight loader.

In darkroom use, simply press the 'RUN' button, unwrap film and feed film into entry slot.

The Velopex processor has a continuous processing mechanism - as soon as the first film has fully entered the mechanism the next film can be inserted.

NOTE: When doing several films it is recommended that the process 'RUN' button be pressed again after the last film has been posted. This avoids the possibility of the machine going into 'standby' while films are still in it.

ENDO INTERRUPT

Processed, but wet, film can be obtained in 2 minutes. By using the endo interrupt facility (see diagram page 3, item 5). To operate, press endo slide downwards to its stop position. After use of the interrupt facility always return the endo slide to its up position for dry-to-dry processing.

IMPORTANT: When using the interrupt facility it is essential to post the film in the centre of the film entry slot, short side leading, using the guide provided. Failure to do so could cause the loss of a film.

Operating the Velopex

FILM COLLECTOR

All films collect neatly in the film collector at the rear of the machine (see diagram, page 3).

EACH DAY

1. Before use check level of developer and fixer and top up where appropriate.

2. Water tank: the water tank on the Intra-X does not have a continuous flow of running water - the water should be replaced with fresh clean water daily.

3. Before processing run through a 'clean-up' film. For clean-up film use a spare intra-oral film which may be kept for this purpose for one week. After a week discard the old film and use a fresh one. This routine helps to clean the transport system and checks that the transport mechanism is running properly.

4. At the end of the day turn the Main Power Switch OFF

WARNING: Turning off electrically at the end of each workday is an essential safety procedure.

# INFORMATION PANEL

## INTRA-X

Serial No. 

WARNING: 
DO NOT USE THE PRESENCE OF 
FLAMMABLE ANAESTHETIC GASES 

EACH DAY 

• Check chemical levels and change water in wash tank.

• Pass cleaning and test film through processor.

• Before processing film check that ready light has gone out.

• After feeding a series of film always re-press button before leaving machine.

• Turn off processor from the mains at night.

WHEN CHANGING CHEMICALS

• Remove wash, fix and developer tanks from unit, empty, clean and dry thoroughly.

• Immerse transport system in hot water (not boiling) and clean with brush provided. Clean all four stages separately in fresh water.

• Refill tanks with fresh chemicals and ensure the transport mechanism is returned to its correct tank as marked.

• Clean body of machine internally and externally.

• Return tanks carefully to their respective positions in the machine and replace panel work.

• Use cleaning film to check correct function.

• Never operate your processor without fluid in all tanks. 

INDICATOR ILLUMINATED - CHEMICAL TEMP. 
LOW INDICATOR OFF - VELOPEX FOR USE

Routine Maintenance

CLEANING THE PROCESSOR

A clean processor is the key to consistent good results and the following cleaning routine should be carried out at each chemical change - the frequency will vary according to the use of the machine and the number of x-ray films passed through it.

NOTE: it is not recommended to extend the use of chemicals beyond three weeks.

CONSULT THE CHEMICAL
MANUFACTURER'S SAFETY SHEETS
BEFORE HANDLING THE CHEMICALS.

1. Open lid using protruding grip, and remove.

2. Remove side cover.

3. Place transport modules in a sink with running water and a drain.

NOTE: the modules will be wet with chemicals - take care to allow them to drain a little over the tanks before moving them.

4. Clean tanks first by rinsing the insides with fresh, running water, and allow to drain. Wipe dry inside and out before re-filling.

5. Clean transport modules thoroughly one at a time, taking care to return each one to the tank from which it has been removed, e.g. Developer module to Developer tank.

6. Carry out the cleaning procedure with fresh water for each module to avoid cross-contamination.

7. The best method of cleaning a module is to immerse it totally in a sink full of hot water (not boiling). Using the brush provided, scrub around the gears and the roller ends.

8. Rinse the belts thoroughly. Using the special tool provided, rotate the gears and belts by hand to assist in thorough cleaning.

9. Do not use detergents -just hot water.

DO NOT USE BOILING WATER AT ANY TIME; BOILING WATER WILL DAMAGE YOUR MODULE.

10. After cleaning, fill the tanks to approximately three-quarters (3/4) of their capacity with fresh chemical, replace the modules in their respective tanks, return them to their correct positions in the machine and top up each one to within 1" (25mm) of the top, taking great care to avoid spillage.

11. The correct order for replacement of tanks and modules is: Developer -Fixer - Water - Dryer.

12. Be sure to keep a note of the date when chemicals are changed, using the chart provided.

Routine Maintenance

PLEASE NOTE: do not begrudge time spent cleaning the machine as you will be amply rewarded by high quality results and extended chemical and machine life.

CAUTION

Processing chemicals should be handled with care. If spilt they may cause staining or corrosion of surfaces. They can also cause skin or eye irritation - wipe up any drips or spillage immediately. If chemicals are splashed onto the skin or eyes, flush immediately with abundant amounts of plain water. Always observe the chemical manufacturer's recommendation for this situation. Problems of this nature can be avoided by wearing protective glasses and gloves.

Troubleshooting


SYMPTOMS AND ACTION

SYMPTOM: Machine does not operate 
ACTION:   1. Check that Mains Power is plugged in and supply switched on.

2. Check by pressing 'RUN' button that machine is not in 'standby' mode.

3. Switch off Mains Power supply and unplug machine. Then check the two fuses in cord socket on the back of the machine. If either fuse is blown, call for Service.

4. The machine is fitted with a safety switch: if the machine lid is not correctly closed the safety switch will prevent the machine from operating. Check by opening and closing again.

5. Check that cord is fully engaged in both plug and socket on the back of the machine.

SYMPTOM: Temperature Indicator Light stays on

ACTION:   (at normal room temperature average warm-up time is 10-15

minutes; in an unheated environment this time could lengthen):

1. If light stays on for an abnormally long time, call for Service.

SYMPTOM: Solution(s) overheating 
ACTION:   Call for Service.

SYMPTOM: Films will not enter through Film Entry Guide 
ACTION:   1. Press 'RUN' button - machine may be in 'standby' mode.

2. Open machine lid and check correct engagement of transport modules.

SYMPTOM: Film Lost in the Machine

ACTION:   1. Check that the transport modules are correctly located in their drive dogs and running correctly. If not, re-locate.

2. Check that endo 'interrupt' bar is fully in the 'up' position.

3. Check the spring tensioning of the modules and check for correct positioning. Use the module turning tool to rotate the belts for inspection, and run a test film through the module using this tool.

4. Check that all transport modules have been put through the correct cleaning procedure.

5. Should there be any damage to the belts, such as tearing or twisting, order a replacement module.

6. Static electricity can retain films in the drying section of the machine. Add a little domestic fabric softener to the washing water when washing dryer module to avoid 'static'.

IMPORTANT: Fabric softener should be used on the dryer module ONLY - do NOT use it on the developer, fixer or wash modules.

Troubleshooting

GEAR REPLACEMENT

If problems are experienced with a transport module the module gears should be inspected for replacement: changing worn gears for new ones is an extremely simple operation and you may choose to do this yourself.

To replace gears, proceed as follows:

Remove retaining screw on gear cover plate (see diagram, item 5);

The gear cover plate (item 6) can now be gently eased off; Remove old gears, and replace with new gears to their correct positions;

Replace gear cover plate and retaining screw.

1.	Main drive gear

2.	'D' shaped centre hole gears (Total 4)

3.	Small idler round hole gear (Total 2)

4.	Large idler gear 
5.	Retaining screw 
6.	Gear cover plate

Film Quality Problems

SYMPTOM: Films too dark

ACTION:  1. Test for light fog by feeding an unexposed film through the machine. It should process as a transparent piece of film base and there should be no shadows or blackness on it.

2. If there is light fog:

i. check that the top lid is firmly in place;

ii. check that the darkroom is light-tight, and that the safelight is sound (e.g. process a test film with safelight off);

iii. check that the daylight loader is firmly secured, including removal of cover tapes from self-adhesive mounting;

iv. check very closely for other possible light leaks: when the daylight loader is used, ensure that the machine has not been sited in direct sunlight or in intense lighting conditions (e.g. directly under striplights or spotlights), do not take your hands out of the daylight loader before the film has fed completely into the machine (fog at one end only of the film indicates premature removal of hands from daylight loader), and check at hand entry ports to ensure there is a good light seal around the wrists - if not, call for service.

v. check whether films have been stored too close to x-ray source, and re-locate.

vi. check expiry date on film box to ensure films are not out of date. (Keep films in a cool, dry place: excessive heat can cause premature ageing of film.)

3. Check that there has been no chemical mix-up, leading to cross-contamination.

4. Check temperature of developer and fixer tanks. These are generally set at: Developer 77°F (25°C), Fixer 82°F (27.5°C). If the Developer temperature is significantly higher, it could lead to dark film. Switch off the machine and call for service.

SYMPTOM: Films too light

ACTION:  1. Check chemicals as follows:

i. chemicals exhausted - replace with fresh;

ii. chemicals contaminated - clean machine, replace chemicals with fresh solutions;

iii. chemical level in tanks too low - top up;

iv. temperature of chemical(s) too low (check with a

thermometer) - generally set at: Developer 77°F (25°C), Fixer 82°F (27.5°C) - if significantly below these temperatures, call for service.

2. Check films have been exposed correctly.

Film Quality Problems

SYMPTOM: Films dirty or marked

ACTION:   1. Check water tank. Clean module thoroughly and change water daily.

2. Check transport modules are being cleaned correctly (see cleaning instructions above).

3. Check chemical level in all tanks and top up as required.

4. Check modules for correct positioning: when found in wrong order, contamination will have occurred. Thoroughly wash and replace with fresh chemicals.

5. Check for stray light entering machine - proceed as for fogging (see above).

SYMPTOM: Films coming out wet:

ACTION:   1. After the processor has been running for 2 mins, check that the air coming out of the vent over the dryer module is warm. If not, switch the processor off at the Mains Power Switch for 2 mins and switch it on again.

2. If the above procedure does not correct the fault, contact your Velopex supplier.

Contamination

When changing chemicals, make sure that the tanks are completely empty and rinsed out. Fresh chemical will be spoiled by contamination, leading to poor results.

WARNING: If the unit should emit abnormal odour, become overheated in some parts or produce unusual noises, immediately switch processor off and unplug from Main Power Supply. Contact your Velopex supplier.

A clean, well-maintained machine makes for good clear x-rays; a dirty machine makes problems.