VELOPEX

SPRINT Mklll

Automatic 
X-ray Film 
Processor

Installation, Operation 
and Maintenance 
manual

WARNING;

Do not plug into the electrical mains power supply before reading this manual - or before filling the machine with liquids

MEDIVANCE INSTRUMENTS LTD

BARRETTS GREEN ROAD, HARLESDEN, LONDON, NW10 7AP
TEL:    +44(0) 20 8965 2913
FAX:    +44(0) 20 8963 1270

Machine serial number to be quoted on all correspondence.

TABLE OF SYMBOLS

| Symbol | #### Description |
| --- | --- |
|  | #### Alternating current |
|  | Off (only for part equipment) |
|  | On (only for part equipment) |
|  | Dangerous voltage |
|  | Caution, hot surface |
|  | Attention, consult accompanying documents |

This processor is fitted with ON/OFF switches and an indicator light illuminated RED. When lit, these indicate that the processor is switched ON and operating correctly. In darkroom conditions, red illumination is used to prevent film fogging and on this processor it is NOT indicting a dangerous condition.









Input rating:

230 Volts, ~ 50 Hz - 200Watt - 0.9 Amps 
This equipment must be plugged into an earthed socket outlet.

Use only the connection lead supplied.

Note: - Replacement fuses must be as follows: - Type, T5AH250V (Time delay 5 Amp high breaking capacity 250 V rated fuse IEC 60127-2 approved).

CONTENTS

Thank you for purchasing a Velopex SPRINT. This model is the simplest and most

compact automatic x-ray processor we produce, and you will have no problem In

finding a convenient site for installation.

Please read this manual carefully before plugging in and using your machine for

the first time. Contained in this manual is information on the right place to position

the Sprint, how to set up the machine correctly for use, and how it should be

maintained so that you receive consistently high quality results over a long

working lifetime.

INTRODUCTION TO THE PROCESSOR

The diagrams at the front of this manual are designed to introduce you visually to

the layout of the Velopex SPRINT.

Page 
TABLE OF SYMBOLS									1 
CONTENTS											2 
TECHNICAL DESCRIPTION								3 
SITING THE SPRINT										4 
EXTERNAL FEATURES OF THE MACHINE						5 
INTERNAL FEATURES OF THE MACHINE 						6 
WALL MOUNTING DIAGRAM								7 
WALL MOUNTING INSTRUCTIONS							8 
PREPARING TO USE THE SPRINT							9 
Filling the Machine with Chemicals								9 
OPERATING THE SPRINT									11 
To Start											11 
Processing Film										11 
Film Collection										11 
Each Day											11 
FILM STORE											12 
ROUTINE MAINTENANCE									13 
Cleaning the Processor									13

Caution											13 
TROUBLESHOOTING									14

Symptoms and Action									14 
FILM QUALITY PROBLEMS								15 
Films too dark											15 
Films too light											15 
Films dirty or marked										16 
CONTAMINATION										16 
USER SERVICEABLE ITEMS								17 
NON USER SERVICEABLE ITEMS								18

TECHNICAL DESCRIPTION

The Velopex "Sprint" X-Ray Processor is a machine for developing and fixing dental x-ray films without user contact between insertion of the exposed film, and receipt of the developed film into the wash-tank for collection at the end of the process.

The machine receives undeveloped film in a light-tight enclosed area, visible through its protective red perspex hood, where a motor-driven module" carries it between two continuous belts through a tank containing Developer. The film is passed out after one minute into an identical "module" working in a second tank, containing Fixer - taking a total of two minutes from film entry to film exit.

Chemical temperature is maintained at 27°C, (81 °F), by two rod heaters protruding from the rear wall of the machine, heating each chemical tank through a metal sleeve.

The film finally passes into a water Wash Tank after processing. It is retrieved by

removing the film from the collector tank.

The mains input socket is fitted with 2x5 Amp fuses, (Type Reference T5AH250V).

The overall dimensions of the processor are:- width 380mm (15 inches), with water tank fitted; height 450mm (17^ inches); depth 230 mm (9 inches).

The weight of the processor is;- empty, 6.85 kg (15 Ib);- full, 10.25 kg (22/z Ib). This machine is a CLASS I device.

ENVIRONMENTAL CONDITIONS FOR TRANSPORT AND STORAGE

This machine should be kept within the temperature range 14°F and 104°F (-10°C and +40°C) and below 80% humidity.

INTERFERENCE BETWEEN EQUIPMENT

This X-Ray processor complies with the following E.M.C. Standards E.E.C- EN60601-1-2 U.L- FCCPart 15 Class A

In the event that electromagnetic or other interference is suspected between this processor and other equipment close to it, move the two units further apart until the problem is eliminated.

SITING THE SPRINT

PRE-INSTALLATION INSTRUCTIONS

Although the Velopex Sprint is designed for use in daylight, care should be taken to avoid sources of intense light. For example do not site directly under a window, fluorescent tube or floodlight. 
A well ventilated position is mandatory.

The SPRINT has been designed for use on a normal counter-top or for

wall-mounting. It is important that when the machine is filled with solution the

stand does not rock or shake, as this could cause spillage. Never move the

machine with chemicals in the tanks.

Avoid siting the machine above or near other electrical, mechanical or sensitive

apparatus - e.g. autoclave, compressor, evacuation equipment.

Place the machine against a wall to prevent accidentally being knocked over.

The machine should not be placed in areas exposed to flammable

gasses.

When the machine is to be sited in a darkroom, it is essential to ensure that the room is well ventilated and light-tight. There should be sufficient space for easy chemical changes and for the machine to be serviced safely. 
When choosing a site for your Velopex it may be convenient to site the machine near a sink for ease of chemical changing.

When choosing the site, it is important to allow for the presence of liquids. Surfaces susceptible to water (or chemical) damage should therefore be avoided. 
Avoid carpeted areas. 
The ambient temperature must be maintained below 27°C (81 °F).

WARNING: X-ray radiation can be harmful to patient, technician and dentist. Inadequate lead shielding of the darkroom or film storage area will also cause fogging of films from exposure to stray x-ray radiation. Consult your Radiation Protection Adviser, your focal bye-laws, your Department of Health and Safety or Dental Equipment Dealer for proper construction of darkroom or placement of film processing equipment in the vicinity of x-ray radiation sources.

CLEANING THE PROCESSOR CASING

Keep the outside of the machine clean with a mild solution of household detergent and warm water (not boiling).

Use a damp cloth only (wring all excess water out of the cloth) and wipe the external surfaces, exercising extreme care around the electrical components.

WARNING: Ensure that the machine is disconnected from the mains power supply during cleaning and never allow water to seep into the casing.

EXTERNAL FEATURES

1 	MAIN CASE

2 	DAYLIGHT LOADER COVER

3	FRONT COVER

4 	HAND ENTRY PORT

5	WASH TANK

6 	FILM ENTRY CHUTE

7 	FRONT COVER LOCATING LUGS

8 	DAYLIGHT LOADER COVER RETAINING CHANNEL

9 	FILM STORE

10 	FILM DISPENSER

11   FILMS

INTERNAL FEATURES 

12 	MAINS SWITCH, FUSED SOCKET & FUSE HOLDER

13 	"RUN" SWITCH, ILLUMINATED

14 	HEATER INDICATOR LIGHT

15 	DEVELOPER TANK

16 	FIXER TANK

17 	DEVELOPER MODULE

18 	FIXER MODULE

19 	FRONT COVER LOCATION SLOTS

20 	DAYLIGHT LOADER

21 	LIQUID LEVEL WINDOW

WALL MOUNTING DIAGRAM

WALL MOUNTING CENTRES

WALL MOUNTING INSTRUCTIONS

The Sprint is Ideal for wall mounting and is fitted with 4 keyhole slots on the rear panel. A further 2 locking holes at the bottom of the rear panel prevent accidental lifting once in place.

Choose an area of wall that Is both flat and upright, ft must be able to carry a load of 25kg, (55 Ib). Use all the mounting points to ensure a secure fixing and to spread the loading onto all the screws.

NOTE: - Please remember to position the processor within 1 metre,

(3 feet), of the electrical supply, preferably with the supply socket to the

right of the machine.

1. Use screws and wall anchors suitable for the type of wall.

2. Using the template provided, hold in the required position on the wall and temporarily fix using suitable adhesive tape. Use a spirit level and adjust as necessary.

3. Using suitable rawlplugs for the type of wall proceed as follows:-

Drill 4 holes in the wall for the keyhole slot screws - To suit No 12 x 35mm (1 1/2 inch) Pan Head slotted screw.

Drill 2 holes in the wall for the lower locking screws - To suit No 8 x 25mm (1 inch) Pan Head slotted screw.

4. Place all the rawlplugs into their appropriate holes.

5. Fit the 4 keyhole screws into position leaving the screw heads 3mm. (3/8 inch), proud of the wall surface.

6. Lift the EMPTY processor up to the wall and using the keyhole slots, hook the machine onto the 4 screws. If necessary, adjust the screw head heights if they appear too tight or loose.

7. IMPORTANT: - Lock the machine into position by assembling the two remaining screws into the lower fixing holes.

8. Test the machine is mounted correctly by checking that it cannot now be lifted off the wall.

PREPARING TO USE THE SPRINT

The machine is supplied in a single carton containing:-

Velopex SPRINT Processor; Operator's Manual; Water tank; Machine Accessory bag

containing; - Electric Lead; Transport Module Turning Tool; Cleaning Brush;

Chemical Change Chart and Warranty Card.

1. Familiarise yourself with the layout of the machine by referring to illustrations at the front of this manual.

2. Remove the water tank and Film accessory bag that are stored inside the daylight loader, (item 20). Entry into the Daylight Loader is made by vertically lifting the daylight loader cover, (item 2).

3. When the machine has been unpacked and in position, the inner packaging must be removed. The transport modules inside are protected by internal packing pieces, open the front cover, (item 3), and discard the packaging along with the separating sheet between the output rollers of both modules, (items 17 and 18).



FILLING THE MACHINE WITH CHEMICALS

WARNING: Ensure the machine is disconnected from the mains power

supply when filling with chemicals and water.

NEVER turn on the machine with the tanks empty.

Hot parts are exposed when removing the tanks. SWITCH OFF THE

MAINS SUPPLY.

Do not use chemistry or film designed for manual processing.

Note: Optimum processor performance will be achieved by using Velopex Fixer and Developer. The machine is built, however, to work satisfactorily using most high quality brands of chemical.

ONLY use chemicals designed for automatic processing. Chemicals used for MANUAL PROCESSING ('hand dunking') will produce poor results.

1. The correct positions for the tanks are clearly marked.

2. Each tank with its transport module in place should be withdrawn:

NOTE; This is done by sliding FORWARDS (towards the operator) until completely clear of its rod heater element fixed to the machine internal back wall. The Developer tank is removed first; the Fixer tank second. Do not lift either tank until it is clear of the front of the machine.

3. Tanks are now free to be taken to a clean working surface for filling.

4. The transport modules should now be lifted from their respective tanks and the separating sheets removed from between the output rollers.

PREPARING TO USE THE SPRINT

5. Start to fill each tank with the appropriate chemical, pouring in 1 litre fluid at this stage. Pour in accordance with the marking on the tanks, i.e. Developer always in the Developer tank and Fixer in the Fix tank.

6. The rod heating elements fixed to the internal wall of the machine must be lightly lubricated only with the silicone grease provided. (Replacement tube Pt. No. I/Fit 5041 P) Use at first time of operation and after every tank removal.

7. Carefully lower each transport module into its respective tank - NOTE, rapid immersion will cause excessive displacement of chemical leading to spillage. Should any spillage occur it should be cleaned up immediately.

8. The level of the liquids should now be adjusted: - top up to within 25mm, (1 inch), from top of the tank, (shown as half way in the tank viewing window, item 21). This level should be checked every day and topped up as necessary.

IMPORTANT: When returning the transport modules to their tanks, ensure that they are correctly positioned, i.e. with the drive pins inwards so that they engage correctly with the drive dogs on the inside wall of the tank area.

9. Return tanks to their individual positions in the machine (clearly marked), -N.B. fixer first, developer second. Push the tanks carefully home until they touch the rear wall of the tank area, ensuring that each drive pin engages its drive dog.

10. Load the film store, (item 9), if required.

11. Replace the front cover by dropping its locating fugs (item 7) into the front cover locating slots (item 19). Now close the front cover until it is clips shut.

WARNING: If the front cover will not close fully, remove ft and make sure that both the tanks touch the rear wall of the tank area and the modules are fully engaged in the drive dogs.

12. Gently replace the red daylight loader cover, (item 2), by lowering it vertically until it engages fully into both front and rear retaining channels (item 8) in the Daylight Loader, (item20).

WARNING: If the red cover is not fully in position the unit will not be light-tight, and film fogging will occur.

13. The wash tank (item 5) should now be filled with water to the "level" line and hung in position on the left hand side of the main case, at the film exit slot.

OPERATING THE SPRINT

## TO START

NOTE: - Ensure that the two tanks are correctly filled with chemicals and

fitted into position before switching on the machine.

First plug the electric power lead into the socket, (item 12), on the right hand side of the processor and plug the other end into the mains wall socket. The processor may now be switched on.

If the chemicals are below operating temperature the Heater Indicator Light, (item 14), will illuminate. This indicates that the SPRINT is bringing the chemicals up to the correct temperature. Switch on the run switch (item 13) until the temperature has been reached, and then switch it off until processing is required. 
Do not begin processing until this light has gone out.

PROCESSING FILM

Switch on the run switch (item 13). As the processor is almost silent in operation,

the run switch is illuminated to indicate that the film processing modules are

running.

Push hands, (with the exposed film package) through the hand entry ports,

(item 4), at each end of the Daylight Loader, (item 20).

Strip the film from its wrapper.

Slide the film carefully down the Film Entry Chute, (item 6) until it is picked up by

the transport mechanism.

As soon as each film has fully entered the transport, you may begin processing the

next.

IMPORTANT; Do not remove hands from the loader until the last film

has entered the transport system. To do so may cause film fogging.

FILM COLLECTION

Films will collect in the water in the Wash Tank and may be lifted clear by hand. Films should be separated before drying to avoid sticking together. The water in the Wash Tank (film collector) should be changed with each processing run of five or more films.

EACH DAY

1. Before use check levels of developer and fixer and top up where appropriate.

2. Before processing run through a 'clean-up' film. This may be done during the warming up period. For clean-up films use a spare intra-oral film which may be kept for this purpose for one week. After a week discard the old film and use a fresh one. This routine helps to clean the transport system and checks that the transport mechanism is running properly.

3. Renew the water in the wash tank.

4. At the end of the day turn the Main Power Switch OFF

WARNING: Turning off electrically at the end of each workday is an essential safety procedure.

FILM STORE

The SPRINT is fitted with a simple film store and dispenser drawer for standard

intra-oral film (item 9).

To load the film store, firmly unclip the top from the front cover and remove.

Stack new film neatly into the back of the film store, (item 11), as illustrated,

uniformly orienting in one direction.

Return the film store back into position and clip the top into place.

Films are simply dispensed by placing the index finger through the U-shaped

dispenser hole (item 10) on the outside of the cabinet, and gently easing them

out.

NOTE: The film store must always be in position, even if empty or not being used, otherwise light will leak into the processing area and will fog the films being processed.

ROUTINE MAINTENANCE

CLEANING THE PROCESSOR MODULES

A clean processor is the key to consistent good

results and the following cleaning routine should

be carried out at each chemical change - the

frequency will vary according to the use of the machine and the number of x-ray

films passed through it.

Regularly inspect the supply cord for any damage, at least every time

the machine is cleaned. Replace if necessary.

NOTE: It is not recommended to extend the use of chemicals beyond three weeks.

1. Switch off the processor and disconnect it from the mains power supply.

2. Place transport modules in a sink with running water and a drain. Rinse under the running water.

Do not allow chemicals to dry on the belts before cleaning. NOTE: the modules will be wet with chemicals - take care to allow them to drain a little over the tanks before moving them.

3. Clean tanks first by rinsing the insides with fresh, running water, and allow to drain. Wipe dry inside and out before re-filling.

4. Clean transport modules thoroughly one at a time with fresh water for each module to avoid cross-contamination. Take care to return each to the tank from which it has been removed, e.g. Developer module to Developer tank.

5. The best method of cleaning a module is to immerse totally in a sink full of hot water (not boiling). Using the brush provided, scrub around the gears and the roller ends. Using the special tool provided, rotate the gears and belts by hand to assist in thorough cleaning. (Rinse the belts thoroughly - Do not use detergents -just hot water.)

Never use boiling water; boiling water will damage your module.

6. After cleaning, pour 1 litre (1 1/4, pints) of fresh chemical into the correct tanks and replace the modules into their respective tanks. Top up the chemicals to within 25 mm (I inch), of the top of the tanks. Renew the silicone grease on the heating elements. Replace the tanks in the processor, fixer first, followed by the developer.

7. Be sure to keep a note of the date when chemicals are changed, using the

chart provided.

PLEASE NOTE: Spend time cleaning the machine as you will be amply rewarded by high quality results and extended chemical and machine life.

CAUTION: Processing chemicals should be handled with care. If spilt they may cause staining or corrosion of surfaces. They can also cause skin or eye irritation - wipe up any drips or spillage immediately. If chemicals are splashed onto the skin or eyes, flush immediately with abundant amounts of plain water. Always observe the chemical manufacturer's recommendation for this situation. Problems of this nature can be avoided by wearing protective glasses and gloves.

TROUBLE SHOOTING

SYMPTOMS AND ACTION

Symptom: Machine does not operate.

Action:    1. Check that Mains Power is plugged in, the supply is switched on and the run switch is in the 'On' position.

2. Switch off Mains Power supply and unplug machine. Then check the two fuses in cord socket on the right hand side of the machine. If either fuse is blown, call for Service.

3. Check that cord is fully engaged in both mains power supply and socket on the right hand side of the machine.

Symptom: Heater Indicator Light Stays On.

Action:     1. If light stays on for an abnormally long time, call for Service. (at normal room temperature average warm-up time is 10-15 minutes; in an unheated environment this time could lengthen).

Symptom: Chemicals(s) overheating. 
Action:    1. Call for Service.

Symptom: Films will not enter through Film Entry Guide. 
Action:    1. Open machine Daylight Loader Cover and Front Cover and check 		  correct engagement of transport modules.

Symptom: Film Lost in the Machine

Action:     1. Check that the transport modules are correctly located in their drive dogs and running correctly. If not, re-locate.

2. Check the spring tensioning of the modules and check for correct positioning. Use the module turning tool to rotate the belts for inspection, and run a test film through the module using this tool.

3. Check that transport modules have been put through the correct cleaning procedure.

4. Should there be any damage to the belts, such as tearing or twisting, order a replacement module.

FILM QUALITY PROBLEMS

SYMPTOMS AND ACTION

Symptom: Films too dark or fogged. 
Action:    1. Test for light fogging by feeding an unexposed film through the

machine. It should process as a transparent piece of film base with

no signs of shadows or blackness evident.

2. If there is light fogging proceed as follows:-

a) Check that the daylight loader cover is firmly in place.

b) Check very closely for other possible light leaks: ensure that

the machine has not been sited in direct sunlight or in intense

lighting conditions (e.g. directly under strip lights or spotlights).

Do not take your hands out of the daylight loader before the

film has fed completely into the machine.

(fog at one end only of the film indicates premature removal

of hands from daylight loader).

Check hand entry ports to ensure there is a good light seal

around the wrists -if not, call for service.

c) Check whether films have been stored too close to x-ray source, and re-locate.

d) Check expiry date on film box to ensure films are not out of date. (Keep films in a cool, dry place: excessive heat can cause premature ageing of film).

3. Check that there has been no chemical mix-up, leading to cross-contamination.

4. Check temperature of developer and fixer tanks. This is generally set at 27"C (81 °F). A substantial difference, could lead to dark film. Switch off the machine and call for service.

Symptom: Films too light.

Action:    1. Check chemicals as follows:

a) Chemicals exhausted - replace with fresh.

b) Chemicals contaminated - clean machine, replace chemicals with fresh solutions.

c) Chemical level in tanks too low - top up.

d) Temperature of chemical (s) too low, (check with a thermometer) - below 27°C (81 °F) - call for service.

2. Check films have been exposed correctly.

FILM QUALITY PROBLEMS

Symptom: Films dirty or marked 
Action:      1. Check water. Clean tank thoroughly and change water.

2. Check transport modules are being cleaned correctly, (see cleaning instructions).

3. Check chemical level in all tanks and top up as required.

4. Check modules for correct positioning. When found in incorrect order, contamination will have occurred. Thoroughly wash and replace with fresh chemicals.

5. Check for stray light entering machine - proceed as for fogging, (see above).

CONTAMINATION

When changing chemicals, make sure that the tanks are emptied completely and rinsed out. Fresh chemical will be spoiled by contamination, leading to poor results.

WARNING: If the unit should emit abnormal odour, become overheated in some parts or produce unusual noises, immediately switch processor off and unplug from Main Power Supply. Contact your Velopex supplier.

A clean, well-maintained machine will produce good clear x-rays, a dirty machine will cause problems.

Do not struggle with film processing problems. Through experience, VELOPEX probably has the answer.

USER SERVICABLE ITEMS

### FUSES

The fuses are located just below mains input socket (item 12) on the right hand side of the processor. For continued protection against fire or electric shock, replace fuses only with the same type and rating as marked on the label adjacent to the fuseholders.

### MODULE GEARS

If problems are experienced with the transport module it should be considered for replacement.

However changing of the gears is an extremely simple operation and can be carried out on site. To replace gears proceed as follows:-

Remove gear cover plate retaining screw, see diagram, (item 22). The gear cover plate, (item 23), can now be gently eased off. Remove old gears and replace with new gears to their correct positions. Replace gear cover plate and retaining screw.