# VELOPEX SYSEMS CLEANER

# INSTRUCTION

# ROUTINE CLEANING

Used for machines which have been regularly cleaned, and are free from heavy deposits on modules or tank walls.

1.  Drain tanks thoroughly, starting with the wash. Refill with cold water and agitate by running the machine until the heater light is extinguished.

2.  Drain again and half fill with cold water. Add 5% of the tank volume (ie 200ml for an Extra-X, 60ml for an Intra-X, 500ml for an MD2000) of machine cleaner to each tank. Top up with cold water. Agitate as above to mix the solutions. Run the machine until the temperature light goes out. Then operate the film entry button to initiate a 15-20 minute cycle.

3. Repeat 1.

4.  Drain and refill with fresh chemicals as normal.

# EXTENDEND CLEANING

Recommended for the first cleaning cycle and every fourth subsequent cycle.

1. Empty tanks, remove modules and wash thoroughly in running water to remove all traces of processing chemical. Use the handbrush supplied with the machine to remove any obvious deposits from the tanks and transport modules. Turn off water supply.

2.  Fill the three tanks approximately two thirds full with cold water. Add approximately 5% of the tank volume (200ml Extra-X, 60ml Intra-X, 500ml MD2000) Velopex Machine Cleaner to each tank. Mix by replacing the transport modules, topping up with water and gently (to avoid spillage) 'dunking' the modules up and down several times (Intra-X), or replacing the lid and switching on (Extra-X, MD 2000).

3.  Run the machine as normal until up to temperature (heater light off). Then operate the film entry button to initiate a 15-20 minute running cycle. At the end of this cycle, remove the wash transport module to sink, and replace with dryer module. Operate for a further processing cycle.

4.  Repeat stage 1 to remove all traces of cleaning chemical from all modules and tanks.

5.  Refill with processing chemicals as per normal chemical change. Turn on water supply.

NOTE:

At the end of the cleaning cycle soaking, the diluted Systems Cleaner in the tanks should be emptied directly to the foul water drain. If used processing solutions are being saved for later collection, DO NOT drain Systems Cleaner into either developer or fixer collection container.

# HINTS AND TIPS

To minimise cleaning and optimise x-ray quality and solution life:

1. The intention is that cleaning is carried out on a regular basis such as every month (depending on usage). If the machine has got a particularly heavy deposit, cleaning may have to be extended and/or repeated.

2.  Changing chemicals before they are excessively exhausted, changing the static wash (Intra-X) frequently and avoiding cross-contamination of processing solutions (e.g. when refilling or topping-up tanks), will help to maintain machines in a clean state.

3.  Maximise the life and color of any unused developer by carefully squeezing the bottle to expel the air above the solution and replacing the cap. A good re-seal is obtained if the inner septum seal has been carefully removed, leaving the rim surface intact.

# HEALTH AND SAFETY

First Aid Action

IRRITANT. KEEP OUT OF REACH OF CHILDREN. The cleaner concentrate contains alkaline sodium hypochlorite (<5%) and cationic detergent as active ingredients. Irritating to eyes and skin. Avoid direct contact and wear rubber gloves. If splashed in eyes, wash out with water and seek medical advice. If swallowed, drink plenty of milk and water and seek medical advice.

This product should always be used in well ventilated surroundings, 10-12 air changes per hour are recommended.

Diluted solution is fully biodegradable in the environment and can be safely disposed of into the sewage system.

Spillage

For small quantities hose into drain with plain cold water or mop up as appropriate. Wear protective clothing and clean mops etc after use. This solution in both concentrated and diluted forms is capable of corroding some metals and damaging absorbent materials such as textiles. Rinsing or soaking immediately with cold water is the best method of avoiding such problems.

WARNING: Avoid contact with acidic materials, such as some toilet and multi-surface cleaners, since this may liberate toxic gas (chlorine).

Storage

Store upright in a cool, well ventilated place, only in the original container. Solutions are not flammable but can produce unpleasant, probably toxic, gases if heated to high temperatures.

Materials Safety Data Sheet (MSDS)

A MSDS is available on application to the Chemical Manager at the address given below.